## Slide 1: Potential Use of EOS-BPM in PWFA Experiments
FACET-II PWFA Meeting
February 2, 2026
Michael Litos
Notes: Done

## Slide 2
EO Cross Polarization Scheme
k
Laser
Polarizer
Horizontally polarized
γ
e- beam
Polarizer
Vertically Polarized
Image on camera
Witness
Drive
CMOS Camera

## Slide 3
EOS-BPM Mk. 1
EOS BPM Mk.1
EO Crystal
Laser
EO Crystal
e- beam
Leverages correlation between strength of phase retardation and location of e- beam with respect to crystal2.

## Slide 4: EOS-BPM Mk. 1 - Signal
Unapertured: Significant clipping on optics  strong diffraction; However, stronger signal/noise won out.
Apertured: Laser was reduced in size to avoid clipping, but signal/noise suffered as a result.

## Slide 5: EOS-BPM Mk. 1 – Laser Timing Stabilization

## Slide 6
EOS BPM Mk.1 – FY25 Experimental Data
EOS2 Signal Strength (arb)
EOS1 Signal Strength (arb)
Δx measured at OTR (μm)
Δx measured at OTR (μm)
Δx from EOS-BPM (μm)
Beam position monitor for high intensity environments
Claire Hansel
65 μm Resolution!
EOS1
EOS2

## Slide 7
EOS BPM Mk.2 Design
4 Crystal Design
Axicon
e- beam
θ
EO Crystal
k
k
Tilted Pulse Front
Axicon creates circular donut
4 crystals around e- beam axis
Easier alignment more uniform laser intensity
More complete azimuthal coverage allows for position measurement in x and y.
Enhanced signal to noise leads to better temporal and transverse resolution.
5o angle of incidence (θ)
3 ps temporal window
4 10mm x 10mm 100μm thick GaP wafers

## Slide 8
EOS-BPM Mk. 2 CAD Model/Pictures
Mounted 100μm Glass Wafers in EOS BPM Mk.2 Crystal Holder
e- beam

## Slide 9
Limiting Aperture
- Crystal minimum aperture is 1 cm
Mirrors have 4 mm hole diameter and are 6.7 mm thick.

## Slide 10: EOS-BPM Mk. 2 Photos
Outside PB
Inside PB
EO Crystals

## Slide 11
EOS BPM Mk.2 Simulation
t
Drive
Witness
Parameter | Value | Unit
Laser Wavelength | 800 | nm
Laser Angle | 5 | degree
GaP Wafer Thickness | 100 | μm
Drive Beam Charge | 1000 | pC
Witness Beam Charge | 800 | pC
Time Delay btw. Beams | 450/135 | fs/μm
Transverse Beam Offset Drive | ~10 | μm
Transverse Beam Offset Witness | ~10 | μm
Bunch Length (Both) | 50/15 | fs/μm

## Slide 12
EOS BPM Mk.2 Time of Arrival Analysis

## Slide 13
EOS BPM Mk.2 Time of Arrival
SNR Drive Beam: 2.20
SNR Witness Beam: 0.92
SNR Drive Beam: 73.1
SNR Witness Beam: 50.4

## Slide 14
EOS BPM Analysis Algorithm
Expected Signal
α1 = α +π/2
Mask to get 1 ring
Piecewise fit to model
Extract r and α to find bunch position
Repeat for 2nd ring
α1
e-beam
α
Crystal rotated 90o
b
a
a
b
α
α1
α
b
a
a
b

## Slide 15: Drive Beam Position Reconstruction
(996 um, -999 um)
Pixels = (2856, 2856)

## Slide 16: Offset Bunch with Different TOAs
Parameters:
1.6 nC
40 um long
Δt: 0.1 ps
T: 2.4 %
Δt: 1 ps
T: 0.53%
Δt: 0.5 ps
T: 1.1%
Δt: 2 ps
T: 0.21%

## Slide 17: On-Axis Two-Bunch Simulation
Parameters:
Drive: 1 nC , Witness: 0.5 nC
30 um bunch length, 400 fs seperation
Δt (laser & ebeam): 0.5 ps
* This assumes t= 0 is the inner edge of the donut beam
Lineout Through Center

## Slide 18
EOS BPM Mk.2 BPM and TOA Results
3 fs relative TOA resolution
<5 μm transverse resolution

## Slide 19: Current Laser Donut Profile
Good enough for commissioning
Need alignment improvement
Apparent spherical aberration  can be corrected in analysis  still want to improve!
Prefer more uniform intensity  background camera coming

## Slide 20: EOS-BPM Summary
EOS-BPM Mk. 1 pretty promising results
EOS-BPM Mk. 2 installed
  Initial configuration strangled signal (too much ND, bad polarization)
  Signal search this week with improved setup
  Science shifts in mid-Feb.
Hope for ~fs TOA and <5 µm BPM capability
  Often >1m upstream of e-beam waist  ~µm effective BPM resolution at waist
Goal: live analysis tool
  Longitudinal bunch separation
  Transverse alignment
  Use in ML algorithms?
Importance to PWFA
  Chromatic emittance growth from transverse offset
  Hosing instability