## Slide 1
Doug Storey
May 3, 2022
Spectrometer electron diagnostics and expected performance

## Slide 2: Current state of the Dump Table
Electron devices (covered in this talk):
  DTOTR1 and DTOTR2 – YAG and OTR
  LFOV – DRZ-FINE + Phosphor
  CHER spectrometer
  PRDMP – mainly used by ops
Gamma screens (covered by Pablo in next talk):
  GAMMA1 – DRZ + CsI array
  GAMMA2/3 – with camembert filter + DRZ fine
DTOTR1/2
LFOV
GAMMA1
CHER
CHER
PRDMP
GAMMA2/3
Electrons
Gammas

## Slide 3: DTOTR
DTOTR1
  High resolution in-vacuum profile monitor
  Used for TCAV measurements, butterfly emittance, etc
    FOV = 7x7 mm
    Resolution 4.5 µm  0.01% energy resolution
DTOTR2
  Larger field of view option as DTOTR1 (for alignment, etc)
  40mm Ø YAG disk, 30 µm pixel size
    FOV = 40x25 mm
  Can view YAG simultaneously with DTOTR1
DTOTR1
DTOTR2
DTOTR1
DTOTR2
Beam Direction

## Slide 4: DTOTR Expected Performance
Single shot emittance measurement
  Expected Resolution: <5% emittance resolution for measuring matched beams at plasma exit
  Limitations: Phase mismatch and existing correlations
Dispersive quad scan emittance measurement
  Provides energy dependency of emittance and measurement of low energy spread beams
  Limitations: requires multi-shot stability
2-D longitudinal phase space
  Will be used with M.L. techniques to train a virtual diagnostic to provide a non-invasive estimate of the LPS for every shot (C. Emma)
DTOTR1: µm, cm
Input: µm cm
Fit: µm cm
DTOTR1 – 2 Bunch w. XTCAV
= 26 kA, = 65 kA
= 17 kA, = 31 kA
Notes:   Butterfly emittance analysis performed on images generated by PIC simulations:
    Gaussian input beams
    10 GeV to 18 GeV PWFA acceleration
    Energy spread of 0.25% at input  2% at output
    PIC simulations performed by Hiroki Fujii (formerly UCLA)

## Slide 5: DTOTR Initial Performance
Dump table beam diagnostics beam commissioning shifts in August, December 2021
Plots show analysis from data collected on DTOTR2 in Aug 2021 run
  Calibration of energy axis on dump diagnostics
  Measurement of energy spectrum
  Butterfly image captured, however in non-ideal beam and optics configuration
Shift plans written up to redo these and other
measurements within next 2 weeks
Energy Spectrum Measurement
Butterfly Pre-processing

## Slide 6: CHER Spectrometer
Specifications:
Images the cherenkov radiation emitted within a gap between 2 silicon wafers
FOV = 4.5 x 12cm
Imaging resolution = ~230 µm  0.4% energy res.
The large FOV makes CHER the main electron energy spectrometer
  This will be key to showing Chan’s first goal of energy depletion of the drive bunch.
Online, but requires calibration, and characterization of the light collection efficiency across the field of view
E200 Piecewise spectral measurement
CHER Optical Charachterization

## Slide 7: COTR observations on the dump and IP cameras
Intermittent flashes of OTR light seen on various S20 screens
Not correlated with charge measurements
Flashes can be suppressed by inserting an earlier screen
IPOTR2 and DTOTR2 shown here, but seen on almost all IP and dump OTR cameras
From scaling studies below – COTR can emit orders of magnitude more light than incoherent OTR, so even a small charge spike can overwhelm the normal OTR signal
No other screens in
S10 screen in
DTOTR2
IPOTR2
IPOTR2
S10 screen in
IPOTR2 screen in

## Slide 8: COTR mitigation
Laser heater – as Claudio covered
  Will inhibit COTR from microbunching altogether
Using blue pass filters on cameras
  Will block longer wavelengths of light that start emitting coherently first
    Filters installed on IPOTR2 camera to test at next opportunity
Color camera
  Can get cool color images so we can see the color content of COTR
    Have the camera, but requires some additional hardware. Will replace USOTR with color camera
OTR rejection techniques:
  Reject all OTR light and image only scintillator light
    Off axis scintillator imaging
    Time delayed scintillator imaging
Current YAG screen setup:
Beam
COTR light
YAG
lens
camera
Viewing at off angle:
Beam
COTR light
YAG
lens
camera
reflection axis
Acquire image after “prompt” OTR light
a)
b)

## Slide 9: Summary and next steps
Summary
Dump table now fully outfitted with first iteration of diagnostics
Experience has been gained over a few shifts, but more dedicated beam commissioning required to test out emittance and LPS measurements
COTR starting to appear  more sophisticated suppressions required
Commissioning procedures prepared for upcoming beam commissioning shifts:
Dump table e- diagnostics:
  https://docs.google.com/document/d/1A3RfoVOopjoxTbNDSk-5PemZUkAwEaoEGOnekJzoszs/edit?usp=sharing
Gamma diagnostics: Pablo is currently preparing
XTCAV 2D LPS:
  https://docs.google.com/document/d/1jKNhruicX-gDcYPo8y_O7JgQAJ5cplsF7GNMh3K3Jt0/edit?usp=sharing