## Slide 1
LCLS-II-HE Cavity and Cryomodule Testing
James Maniscalco | SRF Engineer | SLAC National Accelerator Laboratory
LCWS2023
17 May 2023

## Slide 2: LCLS-II-HE project scope
Slide adapted from G. Hays
Install low-emittance injector and SRF gun for extended hard X-ray performance
Add 23 additional cryomodules (L4 linac) to increase the LCLS-II accelerator energy to 8 GeV.
Install new cryogenic distribution box and transfer line between the cryoplant and the new L4 linac.
New long period Soft X-ray Undulator
Upgrade the LCLS Hard X-ray endstations for MHz beam and data rates

## Slide 3: LCLS-II-HE cavity and cryomodule requirements
Objective: increase beam energy from 4.0 to 8.0 GeV
  23 new cryomodules added to existing 35
Quantity | LCLS-II Cavities | LCLS-II Cavities in HE | HE Cavities in HE | Unit
Number of 1.3 GHz cavities | 280 | 280 | 193† | 
Minimum average cavity Q0 at nominal Eacc | 2.7 | 2.7 | 2.7 | 1010
Vertical test qualifying gradient | 19 |  | 23 | MV/m
Assumed fraction of failed cavities | 6 | 3 | 6 | %
Nominal average operating gradient | 15.7 in L2-3 | 16.9 in L2-3 | 20.8 | MV/m
Vertical test field emission onset min. | 17.5 |  | No FE allowed | MV/m
Maximum CM captured dark current | 10 | 10 | 30 | nA
† includes injector CM (8) and BCC (1)
Bare 1.3 GHz LCLS-II/LCLS-II-HE cavity

## Slide 4: High-Q/High-Gradient R&D
R&D program following LCLS-II project to improve nitrogen doping recipe and surface processing (2018-2020)
  Research collaboration between SLAC, Fermilab, Jefferson Lab, Cornell University
  Increase cavity gradient without degrading Q0
  Improve uniformity of performance with tighter QA
    Challenge: long lag between vendor activities and cavity tests (2-3 month minimum)
Incorporating lessons learned from LCLS-II
  Furnace contamination: recurring issue that led to significant amount of rework
    improved residual gas monitoring needed
  EP temperature: lower & more stable temperatures correlated with higher quench fields
    more thorough measurement of cavity temperature needed
  Flux expulsion: different material batches require different treatment
    pre-determine required heat treatment temperatures by single cell flux expulsion testing
  In-person vendor oversight: necessary solve issues in a timely manner
    be ready to visit cavity supplier on short notice and frequently

## Slide 5: LCLS-II-HE recipe
Additional changes:
  Require continuous RGA spectrum during furnace runs
  Require continuous monitoring of temperatures during electropolishing runs
  Sort cavity half-cell material by required heat treatment temperature
LCLS-II | LCLS-II-HE
Bulk EP High Temperature Furnace Treatment & Doping (2N6) Fine EP | Bulk EP 1 High Temperature Furnace Treatment Bulk EP 2 (last portion cold) Doping (2N0) Fine EP (all cold)

## Slide 6: LCLS-II-HE cavity scope
10 qualification + 168 main production + 24 additional cryomodule cavities ordered from an industrial vendor in Europe, following competitive bid process
  Qualification cavities went into “verification cryomodule” (vCM)
Cavity qualification testing overseen by Cavity Technical Board (CTB)
Manufacture @ vendor
Analysis by cavity technical board
reporting
oversight
Vertical qual.
@ partner labs
Cryomodule
assembly
PASS
Analysis by cavity technical board
FAIL /
MARGINAL
minor issue: light rework at labs
major
issue
CTB: SME’s from SLAC, JLab, and FNAL
Set aside for possible rework

## Slide 7: Vertical qualification testing – requirements
Gradient: Eacc ≥ 23 MV/m
Quality factor: Q0(21 MV/m) ≥ 2.5×1010
  VT stainless steel flange R = 0.8 nΩ
HOM coupling: Qext > 2.7e11
HOM power: PHOM(21 MV/m) < 1.7 W
Field emission: no detectable FE up to quench
Frequency: f = 1300.25±0.10 MHz
Field probe coupling: 2.5×1011 < Qext < 7.0×1011
Multipacting: fully processed before final Q vs. E
7π/9 mode: avoid mode buildup → quick measurements

## Slide 8: Vertical qualification testing – procedures
Cavities shipped by vendor ready for VT
  High-Q coupler: Qext ≈ 1.5×1010
Tested at partner labs after receipt
  Installed on vertical test stand (up to 3 cavities at same time)
    Active pumping at JLab*; not at FNAL
      *tests after Dec. 2022 have no active pumping
  Instrumentation installed
    Radiation monitors
    Four thermometers (minimum)
    Two fluxgate magnetometers (minimum)
  Fast cool to 4.25 K, with Bamb(Tc) ≤ 5 mG
  Cool to testing T = 2.00±0.02 K
Fermilab vertical test system

## Slide 9: Vertical qualification testing – procedures
Calibrate Q0 measurements
  Average of ≥ 5 decay msmts. @ 6-8 MV/m
Measure Q0 vs. Eacc to quench
  1 MV/m increments (or smaller increments as needed)
  No administrative limit on gradient
Watch for FE and MP quenches
  If Q degradation after MP processing, evaluation by CTB
  If FE present and conditions away, final power rise from low field to quench in 1 MV/m increments required for qualification.
Example vertical test data

## Slide 10: Vertical test performance so far – as of May 10, 2023
145 cavities tested, of which:
103 qualify (71%)
  92 as received
  8 after 1 rerinse
  3 after 2+ rerinses
1 “on hold” for marginal Q0 (<1%)
9 disqualified (6%)
  4 with low Eacc
  3 with persistent FE
  1 with low Q0
  1 with HOM scratch
6 in need of mechanical repair (4%)
26 awaiting re-test after high-pressure rinse to mitigate field emission (18%)
All cavities with no FE

## Slide 11: HE cavities vs. LCLS-II: accelerating gradient in VT
Average quench gradient: 22.0±3.9 MV/m
Average quench gradient: 27.2±3.5 MV/m
LCLS-II
HE vCM + production
>5 MV/m improvement
LCLS-II spec: 19 MV/m
LCLS-II-HE spec: 23 MV/m

## Slide 12: HE cavities vs. LCLS-II: quality factor in VT
Average Q0:
3.18±0.44×1010
Average Q0:
3.20±0.37×1010
LCLS-II
HE vCM + production
Only considering cavities that reached at least 21 MV/m
Very few cavities below spec

## Slide 13: Quality factor by heat lot
Material segregation strategy has been successful
  Higher heat treatment lots generally show higher Q0 (better flux expulsion)
  One lot (T107) appears to have been assigned the wrong temperature. Still meets specification (except 1 cavity)
  Mixed 900 °C cavities have higher Q0 than pure cavities at same temp.
  High Q0 also for mixed 900/950 °C cavities with pre-treated subassemblies

## Slide 14: Cryomodule assembly
Cryomodule assembly at partner labs proceeding. So far:
  4* modules completed and tested with production LCLS-II-HE cavities at FNAL
  1 module has undergone initial testing at JLab
Leveraging assembly lines / test facilities / personnel from LCLS-II assembly
* F1.3-21 built with spare cavities from LCLS-II and 2 spare vCM cavities.
FNAL
JLab

## Slide 15: Cavity performance in cryomodules
F1.3-20 (vCM)
  Average usable gradient: 24.1 MV/m
  Usable voltage: 200 MV
  2 K heat load: 119 W (Q0,avg=3.12×1010)
F1.3-21
  Average usable gradient: 25.1 MV/m
  Usable voltage: 201 MV
  2 K heat load: 141 W (Q0,avg=2.62×1010)
plots from S. Posen

## Slide 16: Cavity performance in cryomodules
F1.3-22
  Average usable gradient: 25.1 MV/m
  Usable voltage: 209 MV
  2 K heat load: 116 W (Q0,avg=3.18×1010)
F1.3-23
  Average usable gradient: 25.1 MV/m
  Usable voltage: 209 MV
  2 K heat load: 115 W (Q0,avg=3.14×1010)
plots from A. Cravatta

## Slide 17: Cavity performance in cryomodules
F1.3-24
  Average usable gradient: 22.6 MV/m
  Usable voltage: 187 MV
  2 K heat load: 115 W (Q0,avg=3.27×1010)
J1.3-22 (preliminary)
  Average usable gradient: 22.4 MV/m
  Usable voltage: 186.2 MV
    Expected to increase after MP processing
  Heat load/Q0 measurements pending restart of testing
plots from A. Cravatta

## Slide 18: Outlook
Cavity deliveries to continue through April-May 2023
Additional 24 cavities to come in batches of 12
  Fall 2023
  Spring 2024
Cryomodule assembly and test to continue through mid 2025