## Slide 1: + UED Beamline Design for Optimising Temporal Resolution
Julian McKenzie (on behalf of the RUEDI team)
UED Opportunities for Dynamical Imaging of Materials
7th November 2023 Santa Fe, New Mexico, USA
Relativistic Ultrafast Electron Diffraction & Imaging

## Slide 2: Contents
The RUEDI Facility Proposal
RUEDI Imaging/Microscopy Line
What do we want from a UED beamline?
  UED Beamline Options
  Bunch length control
  Jitter suppression
Photoinjector
  Dark current
Post-sample beamline
  fs temporal diagnostics
  TEM style optics
Summary

## Slide 3: The RUEDI Facility Proposal
A new UK National User Facility at Daresbury Laboratory
Relativistic Ultrafast Electron Diffraction & Imaging
MeV UED and UEM
Time-resolved pump-probe experiments in both real and reciprocal space
With a large variety of pumps and sample environments
Daresbury Laboratory
2020 - Funding bid
2022 - Design stage
2023 - Bid for capital
2025 - Construction
2030 - User facility
  Wide variety of laser wavelengths
  THz, TW, HCF
  (keV ion source)
  Solid
  Liquid cells
  Liquid/gas jets
  Plasmas
  Cryogenic (down to mK level)
Dynamics of Chemical Change
Quantum Materials & Processes
Biosciences
Materials in Extremes
Energy Generation, Conservation, and Storage
To enable a large range of science:

## Slide 4: Facility concept
Microscopy/
Imaging line
Ultrafast diffraction line
S-band photoinjector 4 MeV
10s of nm rad emittance
Current concept is one gun feeding two separately shielded beamlines
Exploring different options!
ps scale
200fC - 20pC
Stroboscopic + Single-shot
10-100s fs scale
Up to 400 fC
Stroboscopic + Single-shot

## Slide 5: Best spatial resolution
Microscopy/
Imaging line
Ultrafast diffraction line
S-band photoinjector
ps scale
200fC - 20pC
Stroboscopic + Single-shot

## Slide 6: Imaging Beamline
Imaging lens system
Series of magnetic lenses
Aiming for 6500x magnification
Detector
16-64 MegaPixel
direct electron
~6.5um pixel size
Dechirper cavity reduces energy spread
(see next slide)
Objective lens

## Slide 7: Imaging Beamline
Objective lens
The key to the whole system
Samples inserted into middle
Multiple pump laser entry ports
Compatible with commercial TEM sample stages (enabling experiments in solid, liquid, cryogenic, heating etc)
Imaging lens system
Series of magnetic lenses
Aiming for 6500x magnification
Detector
16-64 MegaPixel
direct electron
~6.5um pixel size
Dechirper cavity reduces energy spread
(see next slide)
Objective lens
Comparing round lenses (solenoids)
vs quadrupole quintuplets
Quadrupoles have worse aberrations
Round lens limited to 2 MeV beam energy at 1.7 T peak field

## Slide 8: Microscopy Resolution
Blue – pre-dechirper
Yellow – post-dechirper
Resolution limit simplified…
RF dechirper cavity
  Also used to decelerate from 4 to 2 MeV
For 1nm resolution, need <10-4 relative energy spread
8
Also need ~10s nm rad level emittance to be able to focus beam down to ~10µm samples
Stroboscopic mode (10^6 electrons) has potential to reach down to 1nm resolution
Single-shot mode (10^8 electrons (20pC)) has resolution limited to 50 – 100 nm range, and hard to focus beam onto small samples
Cs = Spherical aberration (<10mm @ 2 MeV)
𝛼 = objective aperture collection semi-angle
Cc = Chromatic aberration (<10mm @ 2 MeV)
𝛿𝛾/𝛾 = relative beam energy spread

## Slide 9
Microscopy/
Imaging line
S-band photoinjector
Ultrafast diffraction line
Best temporal resolution
10-100s fs scale
Up to 400 fC
Stroboscopic + Single-shot

## Slide 10: UED beamline design
Other talks this workshop are on optimising brightness
  (guns, cathodes etc)
So this talk will concentrate on minimising time resolution
  (whilst preserving beam brightness!)
  Minimise time resolution:
    Minimise electron bunch length
    Minimise time jitter (between laser pump and electron probe)
  Maximum brightness
    Maximise bunch charge
    Minimise emittance
What do we want?

## Slide 11: UED Beamline Options:
Gun
Sample
Just gun (+ focussing) into sample chamber
Gun only
Limited control over bunch length

## Slide 12: Bunch length control
There are two contributions to R56 (longitudinal dispersion)
Different energy particles having different velocities
Space-charge causes electrons to de-bunch
For 4 MeV it is 10 fs/keV per metre travelled
Electron time of arrival is related to energy:
Can design a beamline to have fully tuneable R56 (to both signs as well)
Sextupoles can be used to correct higher order (T566)
Different energy electrons travel different paths through a magnetic beamline
Velocity term
Path length term
E
t

## Slide 13: UED Beamline Options:
Gun
Sample
Just gun (+ focussing) into sample chamber
Gun only
Limited control over bunch length
Gun
Sample
Buncher
Additional component to provide velocity bunching
(RF cavity, or THz (R&D))
Gun + (RF) buncher
Shorter bunches but increases jitter
80 fs
rms jitter
120 fs
rms jitter
Gun only
Gun + RF buncher
E
t
E
t
Gun only
Buncher flips chirp
Jitter

## Slide 14: UED Beamline Options:
Gun
Sample
Gun
Sample
Just gun (+ focussing) into sample chamber
Magnetic beamlines
Gun only
Examples of potential beamlines with different topologies
Limited control over bunch length
Many possible topologies
- some allow 1 gun to feed multiple beamlines
More flexibility
Bends allow collimation of dark current
Possible to create very short bunches and simultaneously suppress the jitter
Gun
Sample
Gun
Sample
Gun
Sample
Buncher
Additional component to provide velocity bunching
(RF cavity, or THz (R&D))
Gun + (RF) buncher
Shorter bunches but increases jitter

## Slide 15: Jitter suppression
Gun jitters have a correlated
Set magnetic R56 to the opposite of the gun jitter to cancel it
Set electron beam chirp (using space charge) to enable simultaneous compression and jitter cancellation
Experimentally proven by KAERI
Jitter suppression
Bunch compression
*H.W. Kim et al, "Towards jitter-free ultrafast electron diffraction
technology", Nature Photonics, Vol. 14, Pages 245–249,
2020.
Julian McKenzie, UED Opportunities for Dynamical Materials, Santa Fe, 07/11/2023

## Slide 16: UED beamline choice
  Minimise time resolution:
  Maximum brightness
What do we want?
Magnetic option has the potential for achieving both of these simultaneously
And could remove a lot of dark current
All for 10^6 electrons, after collimation
4 dipole arc
RF buncher
Gun only
TBA

## Slide 17: Example optics for a TBA
400 fC from gun
Collimate to 176 fC
  After magnetic arc
  (10^6 electrons)
  Goal charge for single-shot
Doubles the brightness (0.2 )
Reduces the bunch length from 21  14 fs FWHM
No residual dispersion
  (which can dominate bunch length, space-charge induced)
Gun
Sample
(Triple Bend Achromat)
3 fs
rms jitter

## Slide 18: Photoinjector
2.4 cell S-band RF gun
1 kHz repetition rate
Side coupled RF input

## Slide 19: Photoinjector
2.4 instead of 1.4 cell:
  Beam dynamics studies:
    Main benefit of 1.4 cell was shorter bunch length
    But we have a compressor!
    We are down to cathode emittance limit anyway
  Much lower field required (for same energy beam)
  Much lower dark current produced
 | 1.4 Cell | 2.4 Cell
RMS bunch length [fs] | 148.0 | 232.3
RMS transverse emittance [nm rad] | 35.9 | 37.3
Max field on cathode [MV/m] | 114.8 | 68.8
2.4 cell S-band RF gun
1 kHz repetition rate
Side coupled RF input

## Slide 20: Photoinjector
 | 1.4 Cell | 2.4 Cell
RMS bunch length [fs] | 148.0 | 232.3
RMS transverse emittance [nm rad] | 35.9 | 37.3
Max field on cathode [MV/m] | 114.8 | 68.8
Half-cell length optimised so that effect of phase/amplitude jitter the same – so both can be supressed with same R56 beamline
Operating phase
2.4 cell S-band RF gun
1 kHz repetition rate
Side coupled RF input
2.4 instead of 1.4 cell:
  Beam dynamics studies:
    Main benefit of 1.4 cell was shorter bunch length
    But we have a compressor!
    We are down to cathode emittance limit anyway
  Much lower field required (for same energy beam)
  Much lower dark current produced

## Slide 21: Photoinjector
Normal pulse
Overcoupled with step down
 | 1.4 Cell | 2.4 Cell
RMS bunch length [fs] | 148.0 | 232.3
RMS transverse emittance [nm rad] | 35.9 | 37.3
Max field on cathode [MV/m] | 114.8 | 68.8
Short RF pulse via over-coupling, and stepped RF pulse
Reduces RF pulse length by 2 to get lower overall dark current
Half-cell length optimised so that effect of phase/amplitude jitter the same – so both can be supressed with same R56 beamline
Operating phase
2.4 cell S-band RF gun
1 kHz repetition rate
Side coupled RF input
2.4 instead of 1.4 cell:
  Beam dynamics studies:
    Main benefit of 1.4 cell was shorter bunch length
    But we have a compressor!
    We are down to cathode emittance limit anyway
  Much lower field required (for same energy beam)
  Much lower dark current produced

## Slide 22: Post-sample beamline
22
Post-sample lenses similar to a UEM:
  Variable magnification
  Increase camera length
  Correct for pre-sample optics
  Correct for space-charge
Temporal properties:
  Time of arrival
  Bunch length
Diagnostic suite:
  RF Beam Arrival Monitors
  RF TDC
  THz at sample position
    For bunch length and calibration
  THz post-detector
    For shot-by-shot measurements
  Spectrometer
Detector

## Slide 23: Summary
RUEDI is a proposed UK National User Facility
  Ultrafast Electron Diffraction & Imaging
  Large variety of sample environments
  Wide range of laser pumps, including 10fs HCF, THz, TW
  Capital funding decision expected end of this year
UED Beamline Design
  Magnetic beamlines have a lot of potential
  Ultra-short bunches with jitter suppression
  Preserves gun brightness
Importance of post-sample diagnostics and lenses

## Slide 24: Acknowledgements
Nigel Browning
Angus Kirkland
Yoshie Murooka
Jan Nugara
Will Bryan
Simon Maskell
B. Layla Mehdi
Sven Schroeder
Carsten Welsch
Oznur Apsimon
Gwyndaf Evans
Dwayne Miller
Ian Robinson
Jasper Van Thor
Lee Kelsall
Alberto Acuto
Maulik Patel
and more…
Alex Bainbridge
James Bebbington
Rachael Buckley
Jim Clarke
Louise Cowie
Graham Cox
Joe Crone
Mike Ellis
Aaron Farricker
Omair Ghafur
Anthony Gilfellon
Clive Hill
Phil Hornickel
Ben Hounsell
Frank Jackson
James Jones
Nirav Joshi
Daniel Lake
Storm Mathisen
Russell McLean
Boris Militsyn
Bruno Muratori
Tim Noakes
Tom Pacey
Suzanna Percival
Lewis Reid
Mark Roper
Yuri Saveliev
Ben Shepherd
Rob Smith
Calum Tollervey
Andrew Vick
Alan Wheelhouse
Fatih Yaman
and more…
Daresbury Laboratory