## Slide 1: Electro-Optic Sampling Beam Positioning Monitor for Relativistic Electron Beams
North American Particle Accelerator Conference 2025
August 12, 2025
Elena L. Ros1, Michael L. Litos1, Alexander Knetsch2, Brendan O’ Shea2, Claire Hansel1, Daniel Matteo3, Gerdard Andonian3, Mark Hogan2, Robert Ariniello2, Shutang Meng1, Tara Hodgetts3, Valentina Lee1
[1] (CU Boulder, Physics), [2] (SLAC National Accelerator Laboratory), [3] RadiaBeam Technologies
Notes: Done

## Slide 2
U.S. Department of Energy, Office of Science, Office of High Energy Physics, under Award Number DE-SC001796.
National Science Foundation under Award Number PHY-1806053
and Faculty Early Career Development Program Award Number
PHY-2047083.
Our Research is Funded By…

## Slide 3
Motivation
Independent transverse and longitudinal measurement of drive and witness beam for Plasma Wakefield Acceleration (PWFA).
Require precise transverse alignment for emittance preservation.
Require precise longitudinal separation to minimize energy spread and maximize efficiency.
Non-destructive shot by shot measurement
PWFA Simulation
Witness Beam
Drive Beam
z

## Slide 4
EO Crystal
γ
e- beam
Laser
k
Background: EO Sampling
EO Spatial Encoding Scheme
Laser
k
EO Crystal
γ
e- beam
EO Crystal
γ
e- beam
Laser
k
EO Crystal
γ
e- beam
Laser
k
EO Crystal
γ
e- beam
Laser
k
Rotated Polarization
Temporal information mapped to transverse profile of laser1.

## Slide 5
Background: EO Sampling
k
Laser
Polarizer
Horizontally polarized
γ
e- beam
Polarizer
Vertically Polarized
Image on camera
Witness
Drive
CMOS Camera

## Slide 6
Electro Optic Sampling Beam Position Monitor
EOS BPM Mk.1
EO Crystal
Laser
EO Crystal
e- beam
Leverages correlation between strength of phase retardation and location of e- beam with respect to crystal2.

## Slide 7
EOS BPM Mk1. / Experimental Data
Beam position monitor for high intensity environments
Δx from DTOTR2 (μm)
Δx from EOS-BPM (μm)
EOS2 Drive Bunch Signal (arb)
EOS1 Drive Bunch Signal (arb)
Drive Bunch Δx at EOS-BPM inferred from DTOTR (μm)

## Slide 8
EOS BPM Mk.2 Design
4 Crystal Design
Axicon
e- beam
θ
EO Crystal
k
k
Tilted Pulse Front
Axicon creates circular donut
4 crystals around e- beam axis
More complete azimuthal coverage allows for position measurement in x and y.
Increase in total signal leads to better temporal and transverse resolution.
5o angle of incidence (θ)
3 ps temporal window
4 10mm x 10mm 100μm thick GaP wafers

## Slide 9
CAD Model/Pictures Inside Picnic Basket
Mounted 100μm Glass Wafers in EOS BPM Mk.2 Crystal Holder
e- beam
This design fits on the delay stage currently installed for EOS in the picnic basket.

## Slide 10
Limiting Aperture
- Crystal minimum aperture is 1 cm
Mirrors have 4 mm hole diameter and are 6.7 mm thick.
If this is too tight for some applications we would be able to swap out the breadboard on the stage for a different configuration for those experiments.

## Slide 11
Outside of Picnic Basket
Picnic Basket
EOS Cam 1
EOS Cam 2
Waveplate
Polarizer (H)
Lens 1
Lens 2
Lens 3
Lens 41
Lens 51
Lens 42
Lens 52
Polarizer (V)
Polarizer (H)
Plan to install EOS 2 camera to look at either the “negative” signal or unfiltered laser
“negative” signal: overall more signal to use
unfiltered laser: better background correction
x 1.5

## Slide 12
EOS BPM Mk.2 Simulation
Parameter | Value | Unit
Laser Wavelength | 800 | nm
Laser Angle | 5 | degree
Drive Beam Charge | 1000 | pC
Witness Beam Charge | 800 | pC
Time Delay btw. Beams | 450/135 | fs/μm
Transverse Beam Offset Drive | ~10 | μm
Transverse Beam Offset Witness | ~10 | μm
Bunch Length (Both) | 50/15 | fs/μm
GaP Wafer Thickness | 100 | μm
t
Drive
Witness

## Slide 13
EOS BPM Analysis
α
e-beam
Intensity on Camera
Temporal window
θ is laser angle of incidence
r

## Slide 14
EOS BPM Analysis with Rotated Crystals
α1
e-beam
Intensity on Camera
Temporal window
θ is laser angle of incidence
α
α1 = α +π/2
Crystal rotated 90o
Crystal rotated 90o

## Slide 15
EOS BPM Mk.2 Time of Arrival Analysis

## Slide 16
EOS BPM Mk.2 Time of Arrival
SNR Drive Beam: 2.20
SNR Witness Beam: 0.92
SNR Drive Beam: 73.1
SNR Witness Beam: 50.4

## Slide 17
Time of Arrival with 1 crystal
SNR Drive Beam: 1.58
SNR Witness Beam: 0.74
SNR Drive Beam: 11.8
SNR Witness Beam: 4.9

## Slide 18
Study of Longitudinal Separation Resolution
Actual Bunch Separation | Fit
400 fs | 397 fs
405 fs | 402
410 fs | 407
415 fs | 412
420 fs | 418
425 fs | 422
430 fs | 428
~ 3 fs resolution with 15 % noise

## Slide 19
Study of Low Charge Bunches
Bunch Charge (pC) | Fit (fs) | Drive | Witness | Drive | Witness
800 | 447 | 19 | 12 | 450 | 350
700 | 447 | 14 | 9 | 300 | 230
600 | 447 | 11 | 7 | 260 | 200
500 | 447 | 7 | 4 | 190 | 150
400 | 448 | 5 | 3 | 120 | 100
300 | 448 | 3 | 2 | 80 | 70
200 | 450 | 1 | 1 | 50 | 50
100 | 440 | 0.5 | 0.1 | 35 | 35
Lineout SNR
Azimuthal Summing SNR
Actual: 450 fs
5% noise at 800pC

## Slide 20
EOS BPM Mk.2 BPM and TOA Results
3 fs TOA resolution
<5 μm transverse resolution

## Slide 21
Future Outlook
E300/E301: More precise drive-witness beam alignment for emittance preservation and separation for optimal beam loading for PWFA experiments.
E302: Shot-by-shot measurement of the witness offset for hosing studies in PWFA.
E306: Shot-by-shot measurement of witness offset for ion channel laser (ICL) studies.
E308: Witness-drive offset measurement to quantify kick from plasma lens
E320: Improved TOA measurement for laser and e-beam for strong-field QED studies.
E388: Improved drive-witness separation measurement for PWFA chirp studies.

## Slide 22
Summary
Novel ultra-fast optical device for measuring beam centroid location with fs/μm level precision.
  Required for emittance preservation of PWFA.
  Required to minimize energy spread and maximize efficiency.
Expected Resolution:
  Longitudinal: ~ 3fs/1μm for signals with up to 40% noise.
  Transverse: < 5μm for signals with up to 40% noise.
Future Work:
  Install and commission EOS BPM Mk.2 at FACET-II next run.
  Explore the utility of EOS BPM to diagnose tilted beams.

## Slide 23
References
[1] A. L. Cavalieri et al., “Clocking Femtosecond X Rays,” Phys. Rev. Lett., vol. 94, no. 11, Mar. 2005, doi: 10.1103/physrevlett.94.114801.
[2] K. Hunt-Stone, R. Ariniello, C. Doss, V. Lee, and M. Litos, “Electro-optic sampling beam position monitor for relativistic electron beams,” Nuclear Instruments and Methods in Physics Research Section A: Accelerators, Spectrometers, Detectors and Associated Equipment, vol. 999, p. 165210, May 2021, doi: 10.1016/j.nima.2021.165210.
[3] S. Casalbuoni, H. Schlarb, B. Schmidt, P. Schmüser, B. Steffen, and A. Winter, “Numerical studies on the electro-optic detection of femtosecond electron bunches,” Phys. Rev. ST Accel. Beams, vol. 11, no. 7, Jul. 2008, doi: 10.1103/physrevstab.11.072802.
[4] B. Steffen et al., “Electro-optic time profile monitors for femtosecond electron bunches at the soft x-ray free-electron laser FLASH,” Phys. Rev. ST Accel. Beams, vol. 12, no. 3, Mar. 2009, doi: 10.1103/physrevstab.12.032802.

## Slide 24
Questions

## Slide 25
Signal Ringing for Short Bunches
Shorter electron bunches correspond to higher THz frequencies. For GaP the cutoff is ~ 10 THz due to the phase velocity mismatch.
This is better than ZnTe, but there is also the trade off, of lower signal.
The phase mismatch can cause ringing in the signal -> worse TOA resolution.
For FACET-II 100 μm GaP is optimal.

## Slide 26
Example of Ringing

## Slide 27
Single Crystal Signal

## Slide 28
Phase Retardation of Crystal
EO Crystal
e-beam
α
Signal from Phase Retardation

## Slide 29
Phase Retardation of Rotated Crystal Orientation
EO Crystal
e-beam
α
Signal from Phase Retardation
EO Crystal rotated 90o
αr
Ir
Ir
I
I

## Slide 30
BPM Analysis
Region I
Region II
Region I
α2
α1
r2
r1