## Slide 1: Experimental Progress of Passive Plasma Lens at FACET-II
PWFA Meeting
Jul 23, 2025
Presenter: Shutang Meng1
Erik Adli2, Robert Ariniello3, Gevy Cao2, Sebastien Corde4, Christopher Doss5, Keegan Downham3 , Thamine Dalichaouch6 , Claudio Emma3, Spencer Gessner3, Claire Hansel1, Mark Hogan3, C. Joshi6, Alexander Knetsch3, Valentina Lee1, Michael Litos1, Nathan Majernik3, Ken Marsh6, Brendan O'Shea3, Elena Ros1, Doug Storey3, Chaojie Zhang6
1CU Boulder, 2University of Oslo , 3SLAC, 4Ecole Polytechnique, 5LBNL, 6UCLA
Notes: Done

## Slide 2: Thin, Underdense, Passive Plasma Lens (TUPPL)
Thin – PWFA much shorter than one betatron period
Underdense – Nonlinear blowout regime
Passive – No reliance on externally driven current
Plasma Lens – Transverse focusing impulse with negligible energy change

## Slide 3: Attractive Features of TUPPL
Extremely strong focusing
  Orders of magnitude beyond electromagnets, PMQs, APL.
Axisymmetric focusing
  Single lens can achieve symmetric focus in x & y
Ultra-compact
  Plasma lens itself: ~500 µm
  Gas jet & laser hardware: ~1 cm footprint along beam line
Rapidly and easily tunable
  Strength scales with density  gas pressure
  Strength scales with length  laser energy / focus/height above gas jet
Self-aligning
  Central axis of blowout determined by electron beam

## Slide 4: The Importance of Strong Focusing
Matching into plasma stages
  Necessary to prevent chromatic emittance growth
  Quadrupole magnets not strong enough
Divergence control coming out of plasma stages
  Prevent chromatic emittance growth in vacuum from high divergence
  Match injected beams exiting plasma to magnets / undulators
Collider final focus
  Axisymmetric – can reduce length
  Ultra compact and strong – can provide tightest focus
  Serve as proxy for collider FF in strong focusing studies (Oide effect)
Other
  SFQED – increase chi: nonlinear quantum param.
  ICS – increase brightness by reducing source size
  HEDP – increase energy density on target

## Slide 5: Comparison to other focusing optics
PPL focusing strength is orders of magnitude stronger than magnets of equivalent phase advance (normalized length).
Phase advance (normalized length):
Quadruple Magnet
Adapted from Taylor, SLAC-PUB-5621 (1991)
Type | g [kT/m] | K [m-2] | L [mm] | f [cm]
Quadrupole Electro-magnet | 0.01 | 0.3 | 84 | 3990
Permanent Magnetic Quadrupole | 0.5 | 15 | 12 | 564
Active Plasma Lens | 3.6 | 108 | 4.4 | 210
Thin PPL (blowout theory, ) | 1468 | 44000 | 0.22 | 10.4
Thin PPL (June 2025 exp, preliminary) | 437 | 13100 | 0.4 | 19
Type | g [kT/m] | K [m-2] | L [mm] | f [cm]
Quadrupole Electro-magnet | 0.01 | 0.3 | 84 | 3990
Permanent Magnetic Quadrupole | 0.5 | 15 | 12 | 564
Active Plasma Lens | 3.6 | 108 | 4.4 | 210
 | 1468 | 44000 | 0.22 | 10.4
Thin PPL (June 2025 exp, preliminary) | 437 | 13100 | 0.4 | 19

## Slide 6: Experimental Setup
Laser, 3 mJ
E-beam
Chaojie Zhang et al 2021 Plasma Phys. Control. Fusion 63 095011
Top view
E-beam
Limitation: Gas jet “true” backing pressure
5mm ; 50 psi

## Slide 7: Electron bunch Characterization
Requirement:
Long, linearly chirped witness bunch to probe transverse wakefield at multiple longitudinal locations.
TCAV measurement
Example PIC sim
Charge measurement:
Witness: 650 pC
Driver: 780 pC

## Slide 8: Plasma Characterization
No interaction
with interaction
Simulation (Fluid + optical)
Fluent simulation
Radial neutral density profile
assume 1e17
Max width: 132 um
Need to rely on plasma expansion to reach plasma length ~ 500 um.
Split-step Fourier code
Plasma length measurement ~ 400 um
  Scan laser height and observe divergence of the witness beam
Optimize laser timing and height location
  Manual 2D scan to find maximum interactions
  Working point (WP) 1: gas jet at 1.5 cm below e-beam; laser fired 3 ns before e-beam arrival.
  WP2: gas jet at 1 cm below e-beam; laser fired 20 ns before e-beam arrival.

## Slide 9: FACET-II Electron Imaging Spectrometer
Quadrupole magnet triplet and spectrometer dipole magnet
Disperses in y, images in x
Image plane at OTR screen near dump, energy setpoint at 10 GeV
Object plane scanned from the gas jet location to gas jet + 50 cm
M12 scan at the gas jet location
Plasma Lens
OTR Screen
Beam Direction

## Slide 10: Image adjustments and data filtering
Energy centroids jitter; align by the valley location between driver and witness
Filter by BC14 bunch length; still good statistics after filtering.

## Slide 11: Evidence of passive plasma lensing
5 energy slices
Object plane scan with no chromaticity correction
E1  E5 with increasing energy
Very clear upstream shift of and reduction in !
E1
E2
E3
E4
E5

## Slide 12: Spectrometer chromaticity
Example: object plane scan from gas jet location to gas jet + 50 cm.

## Slide 13: Emittance measurement of the witness bunch
Preliminary results
Use 20 energy slices, each about 10 MeV.
Object plane location: PPL exit.
(a) Absolute value comparison
(b) Ratio comparison
Incoming vacuum beam has small normalized emittance in x ~ 10 um.
Significant emittance growth in lens:
  Not in underdense regime
  ≳ 1
Betafunction reduction in lens due to emittance growth.

## Slide 14: Preliminary PIC simulation studies
Simulation parameters:
  Gaussian transverse profile with 20 um spot size and 75 cm beta that match x-plane measurement.
  Current profile from BMAD simulation generated by Claudio, similar to TCAV measurement.
  Plasma density ~ ; plasma length ~ 400 um.
Witness beam samples three regions: first wake, defocusing region, front of the second wake.
Good agreement!
150 um
Simulation value scaled by 6%

## Slide 15: Beam size evolution in vacuum
No chromatic correction
Still working on correction of chromaticity for the object plane scan
Strong focusing of 200 – 300 pC of charges.
from 20 um initially
cm from 75 cm initially.
um
21 cm
um
14 cm
PIC simulation
First wake

## Slide 16: June 2025 Experimental Summary
Evidence of a thin passive plasma lens!
  Plasma length ~ 400 um and density ~ .
  Ultra-strong focusing gradient for charges of several hundred pC.
  β* reduced by greater than 72% from 75 cm to below 21 cm.
  Emittance increase up to 100% (quasi-linear regime, beam spot size larger than wake size)
Need more in-depth analysis and simulations:
  Chromaticity correction for object plane scan
  Use more realistic beam profile for PIC sims
  Simulate plasma expansion from PIC.

## Slide 17: Outlook for next runs
Improve setup:
  Better characterization of the gas jet:
    Know the actual backing pressure
    Be able to tune gas jet pressure remotely
  Transverse propagation of ionization laser with 10 mJ:
    Plasma length controlled by laser focus: short and tunable
    Can go to longer plasma length without relying on plasma expansion
  Higher quality incoming e-beam
    Easier to reach underdense regime
    Allow operation at higher plasma density
Broader parameter scans:
  Better statistics on M12 scan
  Vary density with gas jet pressure
  Vary length with laser properties

## Slide 18: Thank you for your attention!
Notes: Done